Background. Denoising improves fidelity metrics and is widely held thereby to permit dose reduction. It does not follow: a denoiser sees only the image it is given, so hypothesis, input and processed image form a Markov chain, and no test on the processed image dominates the likelihood-ratio test on the input. The bound says nothing about what a given observer recovers, or about exposure.

Purpose. To convert a stated detection requirement into the exposure a low-dose CT protocol needs, by separating the detectability an input offers an ideal observer from the detectability a specified observer achieves, and to use that conversion to test the dose claim made for a denoiser.

Methods. Signal- and background-known-exactly detection of a low-contrast disk: on phantoms, with three classical denoisers and three observers; and on 12 liver cases from LDCT-and-Projection-data, with a synthetic lesion in abdominal soft tissue and residual CNNs trained on 8 cases, evaluated on 4 never-seen cases at two capacities, each paired with an adversarial counterpart. Every arm is referenced to its unprocessed input and its analytic ceiling [1,2], and exposure compared against d' ≥ 5 on a dose simulation rescaling the collection’s measured noise.

Results. The unprocessed input met the requirement down to 0.187 of the routine protocol, so the acquisition already supplied the exposure the task needed; 7 of 8 processed arms needed more than that, the best differing from doing nothing by -2%. Exposure read off a fidelity metric disagreed: the exposure PSNR implied reached 2.04× the one delivered, and 4.8× where d' had already fallen to 1.35. An ideal-observer floor supplied 60% of the exposure required. In the controlled experiments 88.2% of evaluations improved fidelity while detectability fell, denoising helped the inefficient observer more (+1.13), PSNR ranked 9 methods against d' at ρ -0.08, and 0 arms exceeded the ceiling. The adversarial arms showed no detected increase in lesion-shaped response.

Conclusions. A detection requirement converts into an exposure, which is what a protocol is written in. Set the dose from the crossing on the unprocessed input, where the margin was here; do not take an exposure implied by a fidelity metric, which overstated it; and do not use an ideal-observer floor as an operating point, which is not reachable. Clinical use requires scanner-specific calibration and human-reader validation.